## Slide 1: TI measurement techniques for pulsed Lidars – the Current Status

## Slide 2: Outline
Current situation
Recommendations from IEA Task #32 on Lidar
Evaluation on TI measurement techniques : focus on model based filtering
Initial results and end-users feedbacks
Perspectives and next steps

## Slide 3: Lidar principles and metrology
N
E
S
W
u
v
w
Max : 200 m
Min : 40 m
10 minutes wind speed :
-> Over 100 successful verifications with met masts + IEC classifications
Accuracy, precision, repeatability proven and standardized
1 Hz data - turbulence intensity:
Differences with cups known to many extent (IEA expert report) + several useful studies and operational uses in conjunction with small masts

## Slide 4: A bit deeper into turbulence measurement
Lidar vs Sonic
Big
scale
Small
scale
IEA task 32 report on turbulence
IEA task 32 report on turbulence
Lidar measures eddys within the scale of their spatial and time resolutions. Part of small scale eddys are filtered.
In standard dev, calculation of horizontal wind speed, other correlation term may pollute (i.e. <u,w>).
Wind is composed of several frequencies : spatial representation of this frequencies are eddys
+ In complex terrain inhomogeneities will also play a role

## Slide 5: Standard data processing
JUWI\Senvion Presentation
EWEA 2014
ECN report
LAWINE 2014
IFPEN presentation
EWEA 2015
Windcube TI close to sonic anemometer, slightly overestimate cup
12 Windcube datasets analyzed in complex and simple site
Results concluded the Windcube is generally conservative of 0.5-3%
6 different sites
Simple terrain: 5% low WS and 2% high WS
Complex terrain: 10% low WS and 25% high WS
Standard algorithm shows repeatable results and shows agreement with mast in simple terrain. Higher deviations at low speed are observed in complex terrain.

## Slide 6: IEA Task #32 on Lidar
Expert report on Turbulence Intensity measurement by Lidars which include continuous wave and pulsed Lidar.
Technically not-easy to reach but many use cases described
Globally positive toward the capacity of Lidars to measure Turbulence Intensity and provide useful information
Extract of conclusion
“lidars […] do not exhibit any signicant limitation in the technology. […] Some additional tricks […] in either post-processing or scanning configurations are therefore required to obtain meaningful turbulence quantities . “

## Slide 7: Innovative signal processing techniques
Model based optimal
filtering technique ,
IFPEN
VAD or DBS
based derivation
Machine learning
NREL
TRL3/4 : proof of concept
TRL5/6 : Tested in intended environment
TRL9 : technology available
TRL 6/7: Proofs in operationnal environment
SLEMT Model
DTU
+ Improvements shown
- Works on conditions similar to the conditions used for learning the correction
+ Improvements shown
+ Automatic algorithm applicable in all terrain types
- May not fully eliminate the errors
+ Improvements shown
+ « Easy » to apprehend and use
- Correction derived for simple homogeneous terrain only
- Requires atmospheric stability information
+ Widely used in the industry
- Show limited performances in complex terrain and for certain wind conditions
-
Paper published in Wind Energy Science
EWEA 2015
EWEA 2015

## Slide 8: Evaluation of Model-based filtering

## Slide 9: A worldwide trial
From mid-2016, a representative set of Windcube end-users are testing the Model-based filtering technique on their own existing data sets where a met mast is available
Consultant
Turbine manufacturer
Others
Research institute
Developpers, O&M

## Slide 10: Wide spectrum of site conditions
10 months of data were compared to mast inducing consistent statistical representative comparison
Different heights were tested to comprehensively assess the height dependency : 40 meters to 131 meters
Coastal, forested, farmland terrains: all moderately complex.
Climates mild and wet
Mean temperatures from 8° to 16°

## Slide 11: Initial results and end-users feedback

## Slide 12: Results 1/3
Mean deviation with cup anemometer over 6 heights is below 1% of TI
Deviations tends to be de-trended with wind speed
40
60
82
99
120
131
C
C
1,0%
0,8%
0,8%
0,7%
1,0%
0,6%
Mean dev (abs)
Deviations with wind speed at 131 meters

## Slide 13: Results 2/3
Mean deviation with cup anemometer at 100 meters is 1,1% with a correlation of 87%
Terrain is forested : moderately complex
100
C
C
1,1%
Mean dev (abs)
R2 = 87%

## Slide 14: Results 3/3
Mean deviation with cup anemometer over 4 heights is below 0,7%
Terrain is coastal and forested : moderately complex
45
60
75
80
C
-0,7%
Mean dev (abs)
-0,5%
-0,2%
-0%

## Slide 15: Initial results from promising neural network techniques

## Slide 16: Neural Network techniques
Algorithm : machine learning algorithm
Results : sonic and Winducbe
simple terrain
Standard
Neural network
The method requires databases in order to « teach » the neural network
First results show promising enhancement
Y=A*X | Standard | Neural network
Stable | 0,92 | 1,00
Unstable | 1,15 | 1,05
Neutral | 1,10 | 1,07
Comparison mast and Lidar : forced fit

## Slide 17: Perspectives and next steps

## Slide 18: Wrap up
Reasons for differences between mast
and Lidar are known ( see IEA Task #32)
Model-based filtering is deemed to be the most advanced algorithm in terms of industrialization process although interesting techniques are being developed (Neural network, SLEMT)
4 dataset of at least one month has been processed through model based filtering :
  Deviations from mast is less than 1%
  Deviations are de-trended with wind speed
“The WiSE algorithm has clearly made an excellent improvement to the Windcube measure of turbulence intensity in comparison to a cup anemometer” One of the End user

## Slide 19: Next steps
Test will be conducted till end of 2016 with more than 10 participants involved and nearly 30 datasets to be analyzed.
We are open to welcome new participants.
Comprehensive report on performances will be issued and publicly available end of 2016.
Aim : Turbulence intensity measured by Lidar to reach stage 1